## Slide 1: Fast Fourier Transforms (FFTs) and Windowing TIPL 4302 TI Precision Labs – ADCs
Created by Luis Chioye, Art Kay
Presented by Peggy Liska
Notes: Hello, and welcome to the TI Precision Lab covering frequency domain analysis for data converters. This video will introduce the fast fourier transform, or FFT. The goal here is to provide an intuitive understanding of the FFT and the deep mathematical background is intentionally omitted. The concept of windowing the signal to minimize error sources is also introduced.

## Slide 2: Definition for time to frequency transformations
Notes: Here we show the mathematical relationship for the fourier transform and the fast fourier transform. The fourier transform is a general method for taking a time domain equation and converting it to a frequency domain equation. The fast fourier transform, or FFT, is designed to efficiently take discrete time domain data and convert it to frequency domain data. Since the FFT is designed to be used with discrete time data, it is commonly used to convert ADC data from time domain to frequency domain. We will not dig deeply into the math here as there are entire text books on the subject. Suffice it to say that for “N” number of time domain points the FFT transform will produce “N” number of frequency domain points. In the next few slides we will cover some basic concepts with FFTs to give a better intuitive understanding of the subject.

## Slide 3: FFT Basics: Alias and Frequency Resolution
FFT assumes time domain continues forever
Number of points in time domain equals number of points in FFT
The alias region is normally hidden. Mirror image about fs / 2 “aliasing”
Frequency Resolution = Δf = fs / N
Δf = fs / N = 1MHz / 16 = 62kHz
Notes: Here we show the time domain and its associated FFT or frequency domain representation. First, the mathematics behind the FFT assumes that the time domain signal continues to infinite time. The time domain signal at the bottom contains 16 samples. The FFT assumes that these 16 samples are infinitely repeated. Second the number of points in the FFT are equal to the number in the time domain. In this example the time domain has 16 samples, so the frequency domain also has 16 samples. Third, the FFT contains a region called the alias. The alias is a mirror image of the valid frequency region and it occurs above the sampling frequency divided by two. In this example the sampling frequency is 1MHz so everything above 500kHz is an alias. The alias region is normally hidden as it is redundant information. We will discuss aliases more in detail later. Forth, all FFTs have a frequency resolution that it equal to the sampling frequency divided by the number of samples The frequency resolution is the minimum change in frequency that the FFT can detect In this example the frequency resolution is 1MHz / 16 = 62kHz. Thus the measured frequency bins step in increments of 62kHz for this example.

## Slide 4: Alias is a Mirror Image of Sampled Signal
Example of a “square wave”
Square wave contains odd harmonics e.g. 3, 5, 7 …
Notice the symmetry around 500kHz
Notes: This slide is to emphasize that the alias region of the FFT is the mirror image of the signal frequency region. In this example, a square input wave is applied. This square wave has some odd harmonics. Notice how the mirror image of the frequency response shows up in the alias region. Again, this region is often hidden in many different FFT software displays as the information is redundant.

## Slide 5: FFT Example Calculation
Example FFT | 
 | Sampling Rate
 | Number of Samples
 | Frequency Resolution
 | Sampling time
 | Input signal
 | Frequency Bin Note: fin is an exact integer multiple of Δf
Example FFT | 
 | Sampling Rate
 | Number of Samples
 | Frequency Resolution
 | Sampling time
 | Input signal
 | Frequency Bin Note: fin is an exact integer multiple of Δf
Notes: Let’s step through some basic math for an example FFT. First, notice that we have a 1MHz data converter sampling rate. Next we collect a total of 16 digitized time domain samples. The frequency resolution for this FFT is 1Msps/16 or 62.5kHz. The time required for each sample is calculated as one divided by the sampling frequency or 1/1MHz = 1us. For this example the time domain input signal is a 125kHz sinusoidal waveform. Translating the time domain sine wave to frequency domain will create a component in a frequency bin. The frequency bin is determined by dividing the input signal frequency by the frequency resolution, which in this case is 125kHz/62.5kHz = 2.0. Note that in this case the frequency bin is an integer because the input frequency is an exact integer multiple of the frequency resolution. So, the sampled signal generates a frequency component in the second bin at exactly 125kHz. The math for the translation works out nicely in this example as the frequency bin is exactly two, but what happens when the bin isn’t an integer.

## Slide 6: FFT – Different Input Frequency
Notes: For this example we increased the input frequency to 187.5kHz. Again, this frequency happens to be an exact integer multiple of the frequency resolution so it creates a frequency component in bin 3. However, what happens if the input frequency is not an exact multiple of the frequency resolution?

## Slide 7: FFT – Spectral Leakage
Example FFT | 
 | Sampling Rate
 | Number of Samples
 | Frequency Resolution
 | Sampling time
 | Input signal
 | Frequency Bin Note: fin is NOT an exact integer multiple of Δf
Example FFT | 
 | Sampling Rate
 | Number of Samples
 | Frequency Resolution
 | Sampling time
 | Input signal
 | Frequency Bin Note: fin is NOT an exact integer multiple of Δf
Notes: This slide shows what happens when the input frequency is not an exact multiple of the sampling frequency. In this case all the math is the same except that the input frequency was changed from 125kHz to 180kHz. Calculating the frequency bin for 180kHz gives a frequency bin of 2.88. So it’s close to 3 but not exactly 3. This results in smearing the frequency components into multiple bins. The tallest frequency component is located at bin 3 and other frequency components show up in adjacent bins. This smearing of the frequency domain signal is called “spectral leakage”. Spectral leakage is problematic as it is indistinguishable from noise near the signal.

## Slide 8: Window: Eliminates discontinuity in sampled waves
Notes: Remember, the FFT assumes that the time domain signal continues infinitely in time. However, practical real world signals always contain a finite number of points called a record. The FFT copies and repeats each record infinitely. If the time domain signal is an integer multiple of the sampling frequency than each record will connect smoothly to the next record. On the other hand, when the signal is not an integer multiple of the sampling frequency the repeated waveform will not smoothly connect to the previous waveform and there will be a discontinuity. The time domain waveform at the top shows the discontinuity between each record. This discontinuity is what creates spectral leakage. One approach to minimize the spectral leakage is to use a time domain windowing function. A window shapes the time domain signal to force the time domain signal to zero at the end of each record. The figure at the bottom shows the effect windowing has on the figure above. Notice that the discontinuity is masked because the window forces the function to zero at the discontinuity. The objective of windowing is to minimize spectral leakage.

## Slide 9: Comparing Frequency Response of Different Windows
Ideally we would like a very narrow main lobe and very deep attenuation in side lobes.
For ADC characterization 7 term Blackman Harris is most often used.
Notes: The window has a frequency response and function like a band pass filter. The main lobe in the frequency domain window passes the fundamental and the side lobes attenuate the spectral leakage. This graph shows the frequency response for several different filters. Ideally, the main lobe, analogous to the pass band, would be very narrow and the side lobes, analogous to the stop band, would have very high attenuation. In practice, windows with good side lobe attenuation will have wider main lobes. For the purpose of ADC analysis, the wide main lobe isn’t a significant problem, so the Seven-term Blackman Harris window is used to achieve the lease amount of spectral leakage.

## Slide 10: Different Windows for Different Applications
Signal Content | Window
ADC characterization | Blackman Harris
Sine wave or combination of waves | Hann
Sine Wave (amplitude accuracy is important) | Flat Top
Narrowband random signal (vibration data) | Hann
Broadband random (white noise) | Uniform
Closely spaced sine waves | Uniform, Hamming
Excitation signals (hammer blow) | Force
Response signals | Exponential
Unknown content | Hann
Two tones with frequencies close but amplitudes very different | Kaiser-Bessel
Two tones with frequencies close and almost equal amplitudes | Uniform
Accuracy single tone amplitude measurements | Flat Top
Notes: FFT windowing is not just used for ADC characterization but is used in many forms of analysis. Depending on the type of analysis required different types of windows may be required. This table lists different types of windows and an example of what type of signal content this window is optimal for. For example, a uniform window is optimal for measuring white noise. In general, we will be using the seven term Blackman Harris as we are using the FFT for ADC characterization.
Understanding FFTs and Windowing
Publish Date: May 04, 2015,

## Slide 11: Window Processing Errors
Parameters used to characterize frequency response of windowing functions |  |  |  |  |  | 
Window | Highest side lobe level (dB) | Processing loss (dB) | Scalloping loss (dB) | Worst-case processing loss (dB) | 6-dB bandwidth (bins) | Half main lobe width (bins)
No window | -13 | 0.00 | 3.92 | 3.92 | 1.21 | 1
Hann | -32 | 1.76 | 1.33 | 3.09 | 2.00 | 2
Hamming | -43 | 1.34 | 1.76 | 3.10 | 1.82 | 2
Four-term Blackmon Harris | -92 | 3.00 | 0.83 | 3.83 | 2.72 | 4
Seven-term Blackmon Harris | -163 | 4.20 | 0.46 | 4.66 | 3.52 | 7
Notes: Since the window function modifies the amplitude and frequency response of the sampled data there are some associated processing errors. This table summarizes these errors. In most cases these error terms can be accounted for and corrected in software. This is taken care of automatically in our evaluation module software.
Let’s look at each error source. Processing loss is the reduction in SNR due to the signal spreading across the main lobe width. Because of this spreading in the signal power, a gain correction factor must be used to accurately compute SNR. The processing loss assumes the frequency of the test signal falls exactly in the middle of the frequency bin. Since the input frequency may not match exactly the frequency bin in a non-coherent test setup, the scalloping loss is the perceived signal attenuation in the FFT when the signal frequency component falls exactly between bins. The worst-case processing loss is the sum of the processing loss and scalloping loss, and it provides a measure of the maximum SNR reduction due to the window function and the worst-case frequency location.
http://electronicdesign.com/analog/choose-right-fft-window-function-when-evaluating-precision-adcs

## Slide 12: Thanks for your time! Please try the quiz.
Notes: That concludes this video – thank you for watching! Please try the quiz to check your understanding of this video’s content.

## Slide 13: References
Luis Chioye, “Choose The Right FFT Window Function When Evaluating
Precision ADCs”, Electronic Design, November 2013
http://www.electronicdesign.com/analog/choose-right-fft-window-function-when-evaluating-precision-adcs
National Instruments, Understanding FFTs and Windowing, May 2015, http://www.ni.com/white-paper/4844/en/
Bores Signal Processing, “FFT Window Functions: Limits on FFT Analysis,” http://www.bores.com

## Slide 14
© Copyright 2017 Texas Instruments Incorporated. All rights reserved.
This material is provided strictly “as-is,” for informational purposes only, and without any warranty.
Use of this material is subject to TI’s Terms of Use, viewable at TI.com