## Slide 1: Coherent Sampling and Filtering to Improve SNR and THD TIPL 4303 TI Precision Labs – ADCs
Created by Art Kay, Luis Chioye
Presented by Peggy Liska
Notes: Hello, and welcome to the TI Precision Lab covering ways to improve SNR and THD measurements in the characterization of data converters. This video introduces a measurement technique called “Coherent Sampling”. This technique is a method that can be used to get the best possible ADC characterization results. In other words, coherent sampling eliminates error sources related to the discrete nature of the FFT and allows characterization of high performance data converters. Also covered in this lab is a test method that uses an external filter to reduce noise and distortion on the input signal to a level where it allows for accurate characterization of the data converter.

## Slide 2: Ideal Situation: Coherent Sampling
Example FFT | 
 | Sampling Rate
 | Number of Samples
 | Frequency Resolution
 | Input signal
 | Frequency Bin Note: fin is an exact integer multiple of Δf
Example FFT | 
 | Sampling Rate
 | Number of Samples
 | Frequency Resolution
 | Input signal
 | Frequency Bin Note: fin is an exact integer multiple of Δf
Notes: Here we go through an example of a data converter characterization system. Let’s step through the math as we did in the TI Precision Labs video titled “Fast Fourier Transforms (FFTs) and Windowing”. Here the sampling rate is 1Msps and 256 samples are taken. This information is used to find the frequency resolution. The input frequency is 62.5kHz, so dividing the input frequency by the frequency resolution shows that the frequency bin is the exact integer 16. Since the input frequency is an exact integer multiple of the frequency resolution we have no spectral leakage. This is called coherent sampling. Based on this information it seems like a good ADC characterization method is to adjust the input frequency to an integer multiple of the frequency resolution so that we have coherent sampling. Unfortunately that isn’t generally so practical as we will see in the next slide.

## Slide 3: Reality: Non-Coherent Sampling
Example FFT | 
 | Sampling Rate
 | Number of Samples
 | Frequency Resolution
 | Input signal
 | Frequency Bin Note: fin is NOT an exact integer multiple of Δf
Example FFT | 
 | Sampling Rate
 | Number of Samples
 | Frequency Resolution
 | Input signal
 | Frequency Bin Note: fin is NOT an exact integer multiple of Δf
Notes: The reality of most equipment is that the signal generator output frequency and microcontroller clock frequency do not have the accuracy and stability to facilitate coherent sampling. Notice that the sine wave generator is adjusted to 62.5kHz but the actual output frequency is 62.503kHz. Furthermore, the output frequency of this generator drifts with time and temperature. A similar phenomena exists with the microcontroller clock rate. Ideally, the sampling rate is set to 1Msps, but in reality it is at 0.9997MHz and will drift with time and temperature. The end result when you work through the math is that we no longer have coherent sampling, and the FFT will have spectral leakage. So, this simple method cannot be used for coherent sampling. In the next slide we will see how a more sophisticated system can be used to achieve coherent sampling.

## Slide 4: Practical method for coherent sampling
Example FFT | 
 | Sampling Rate
 | Number of Samples
 | Frequency Resolution
 | Input signal Related to fs
 | Frequency Bin Note: fin is an exact integer multiple of Δf
Example FFT | 
 | Sampling Rate
 | Number of Samples
 | Frequency Resolution
 | Input signal Related to fs
 | Frequency Bin Note: fin is an exact integer multiple of Δf
Notes: One way to achieve coherent sampling in a practical system is to synchronize the frequency of the microcontroller with the signal generator. Thus the error in the microcontroller sampling rate will track and cancel the error in the function generator so that the input frequency will always remain an integer multiple of the frequency resolution even when the clock drifts. So, the clock from the microcontroller is also applied to a sync input on the signal generator. Even though synchronizing the signal generator to the microcontroller theoretically solves the coherent sampling issue, high precision PLL clock generators and high end signal generators are typically used in this kind of characterization as clock jitter and drift will introduce other additional error. Also note that this kind of characterization is normally required for high resolution converters with resolution of 18 bits or higher.

## Slide 5: Using a filter to clean up signal source
Notes: Another important consideration in ADC characterization is the purity of the input signal. Many common signal generators have resolutions of 12 bits or less. These devices do not have the THD and SNR capable of testing high performance SAR converters. Some high end signal generators have very good THD and SNR, but these can be expensive and may not be available in your lab. One approach to getting high fidelity signals from a common signal generator is to use a band pass filter at the output of the signal generator. This example figure shows that the output of the signal generator has a relatively high noise floor and significant harmonics. Notice that the function generator is adjusted to generate a 1kHz sine wave. This sine wave is applied to a band pass filter with a center frequency of 1kHz. The pass band on the filter will pass the fundamental and the stop bands will significantly attenuate the noise and harmonics. Normally a high order filter is used to maximize the attenuation of noise and harmonics. Using this method, it is possible to generate an input signal sufficient for testing a 18 to 20 bit converter using a common signal generator.

## Slide 6: A practical example
Number of samples in FFT. Increasing number of samples:
Increases frequency resolution
Minimizes impact of spectral leakage
Many samples increases measurement time and potential of frequency drift
Window type: 7 term Blackman-Harris works best for most ADC characterization.
Notes: This slide shows an example of the FFT feature on our evaluation modules. Notice that SNR, THD, SINAD and other ac performance metrics are displayed with the FFT. Also, you can zoom in on the first 9 harmonics as shown here. The number of samples can be adjusted. Generally, increasing the number of samples will reduce the impact of spectral leakage. However, taking more samples will increase the measurement time. For very long measurement times the signal generator output frequency may drift which will have an impact on SNR. For best performance 16k to 32k samples, is often used, however, this number may be different depending on the device and on the test equipment used. Also notice that the type of window can be adjusted. In general, the 7 term Blackman-Harris is a the best choice.

## Slide 7: Thanks for your time! Please try the quiz.
Notes: That concludes this video – thank you for watching! Please try the quiz to check your understanding of this video’s content.