## Legend
Color Legend
Key Parameter
Insert parameter
Insert Key Parameter
<value> | Calculated Value - Conditions Satisfied
<value> | Calculated Value - Conditions NOT Satisfied
<value> | Calculated Value
Temperature Dependant Parameter/Input
Calculators
Bits to PPM to %FS
32 | [Bits]
0.00023283 | [ppm]
2.328e-08% | [%]
% to ppm
0.005% | [%]
50 | [ppm]
ppm to %
300 | [ppm]
0.03% | [%]

## TIPD188 Total Error Analysis
ADS1262 PARAMETERS
AVDD | 5 | (V)
AVSS | 0 | (V)
PGA GAIN | 32 | (V/V)
RBRIDE NOMINAL | 350 | (Ω)
VREF | 5 | (V)
ADC Noise Resolution
Δ VIN | 0.01 | (V)
Full-Scale Range Utilization | 6.4% | (%)
Data Rate | 20 SPS
Filter | FIR
ADC Noise RTI | 198 | nVP-P
Weight Resolution (DC Excitation) | 15.6241 | Bits
| 50505.0505 | Counts
Weight Resolution (AC Excitation) | 16.1241 | Bits | NOTE: Noise reduced by sqrt(2) to account for averaging
| 71424.9274 | Counts
Error Calculations
System Temperature Range | 165 | (Δ°C)
Input Utilization | 6.4% | (%) | [8 empty] | Error Type |  | TYP
Errors | [10 empty] | NOTES: | - Unknown MAX errors are estimated
FSR | 0.3125 | (V) |  | Before Calibraion |  | After Calibraion |  | Temp Comp. | 10 |  |  | - Load cell errors (such as linearity and self-heating) were not included in this error analysis
Noise | 198 | (nVP-P) |  | 0.6336 | (ppm) | 0.6336 | (ppm) | 0.6336 | (ppm)
Offset (Before Cal) | 10.9375 | (uV) |  | 35 | (ppm) | 0.1584 | (ppm) | 0.1584 | (ppm) | ← Offset Error
(After Cal) | 49.5 | (nVP-P)
Offset Drift | 10.9375 | (nV/ᵒC) |  | 5.775 | (ppm) | 5.775 | (ppm) | 0.35 | (ppm) | ← Offset Error (Temp-Dependent)
Gain Error (Before Cal) | 0.005% | (%) |  | 1 | (ppm) | 0.010138 | (ppm) | 0.010138 | (ppm) | ← Gain error (VIN dependant)
(After Cal) | 49.5 | (nVP-P)
Gain Error Drift | 0.5 | (ppm/ᵒC) |  | 5.28 | (ppm) | 5.28 | (ppm) | 0.32 | (ppm) | ← Gain error (VIN dependant)
INL | 3 | (ppm) |  | 3 | (ppm) | 3 | (ppm) | 3 | (ppm)
IREF Absolute Bias Current | 150 | (nA)
IREF Bias Current Offset | 200 | (nA) |  | 0 | (ppm) | 0 | (ppm) | 0 | (ppm) | ← Gain Error
(REF Input Filter R) | 0 | (Ω)
REF Resistor Mismatch Offset | 0 | (nV) |  | 0 | (ppm) | 0 | (ppm) | 0 | (ppm) | ← Gain Error (Resistor/Current-Dependent)
(R Tolerance) | 1% | (%)
REF Resistor Mismatch Offset Drift | 0 | (nV/ᵒC) |  | 0 | (ppm) | 0 | (ppm) | 0 | (ppm) | ← Gain Error (Resistance-Temp-Dependent)
(Temp Coeff.) | 50 | (ppm/ᵒC)
IREF Bias Offset Current Drift | 0.30303 | (nA/ᵒC) |  | 0 | (ppm) | 0 | (ppm) | 0 | (ppm) | ← Gain Error (Current-Temp-Dependent)
IINPUT Current Error is Insignificant
TOTAL (Estimated) DC Excitation Error |  |  |  | 36.0088 | (ppm) | 8.4057 | (ppm) | 3.1067 | (ppm)
|  |  |  | 11252.7361 | nV | 2626.7789 | nV | 970.8416 | nV
|  |  |  | 14.8199 | (nV/ᵒC) | 14.8199 | (nV/ᵒC) | 0.89817 | (nV/ᵒC)
TOTAL (Estimated) AC Excitation Error |  |  |  | 6.1708 | (ppm) | 6.0893 | (ppm) | 3.0501 | (ppm)
|  |  |  | 1928.3839 | nV | 1902.8973 | nV | 953.1623 | nV
|  |  |  | 10 | (nV/ᵒC) | 10 | (nV/ᵒC) | 0.60606 | (nV/ᵒC)

## ADS126x Noise Table
FILTER SELECTION |  | Noise (uVrms) |  | GAIN | [5 empty] | Data Rate (SPS) | Equivalent Brickwall Filter* Bandwidth (Hz)
Data Rate | Filter | DATA RATE | FILTER | 1 | 2 | 4 | 8 | 16 | 32 |  | SINC1 | SINC2 | SINC3 | SINC4 | SINC5 | FIR
2.5 SPS | FIR | 2.5 SPS | FIR | 0.145 | 0.071 | 0.038 | 0.023 | 0.014 | 0.011 | 2.5 | 5.21 | 1.25 | 0.96 | 0.83 | - | 6.28
5 SPS | SINC1 | 2.5 SPS | SINC1 | 0.121 | 0.058 | 0.033 | 0.018 | 0.012 | 0.008 | 5 | 9.73 | 2.5 | 1.92 | 1.67 | - | 10.72
10 SPS | SINC2 | 2.5 SPS | SINC2 | 0.101 | 0.055 | 0.025 | 0.015 | 0.01 | 0.007 | 10 | 18.05 | 5 | 3.85 | 3.33 | - | 17.79
16.6 SPS | SINC3 | 2.5 SPS | SINC3 | 0.08 | 0.046 | 0.026 | 0.013 | 0.008 | 0.006 | 16.6667 | 28.37 | 8.33 | 6.41 | 5.56 | - | -
20 SPS | SINC4 | 2.5 SPS | SINC4 | 0.08 | 0.043 | 0.02 | 0.013 | 0.008 | 0.007 | 20 | 33.3 | 10 | 7.69 | 6.67 | - | 28.4
50 SPS | SINC5 | 5 SPS | FIR | 0.206 | 0.098 | 0.054 | 0.028 | 0.02 | 0.015 | 50 | 73.96 | 24.96 | 19.23 | 16.67 | - | -
60 SPS |  | 5 SPS | SINC1 | 0.161 | 0.09 | 0.047 | 0.026 | 0.017 | 0.012 | 60 | 86.54 | 29.95 | 23.08 | 20 | - | -
100 SPS |  | 5 SPS | SINC2 | 0.146 | 0.069 | 0.038 | 0.021 | 0.013 | 0.011 | 100 | 133.88 | 49.86 | 38.46 | 33.33 | - | -
400 SPS |  | 5 SPS | SINC3 | 0.128 | 0.067 | 0.033 | 0.019 | 0.012 | 0.01 | 400 | 423.16 | 197.75 | 153.79 | 133.32 | - | -
1200 SPS |  | 5 SPS | SINC4 | 0.122 | 0.063 | 0.03 | 0.017 | 0.011 | 0.008 | 1200 | 1002.1 | 579.71 | 459.94 | 399.66 | - | -
2400 SPS |  | 10 SPS | FIR | 0.284 | 0.142 | 0.077 | 0.041 | 0.027 | 0.023 | 2400 | 1666 | 1118.9 | 912.02 | 797.32 | - | -
4800 SPS |  | 10 SPS | SINC1 | 0.229 | 0.123 | 0.06 | 0.035 | 0.023 | 0.018 | 4800 | 2647.51 | 2075.41 | 1773.81 | 1579.25 | - | -
7200 SPS |  | 10 SPS | SINC2 | 0.193 | 0.093 | 0.048 | 0.028 | 0.019 | 0.016 | 7200 | 3350.48 | 2869.69 | 2557.96 | 2332.4 | - | -
14400 SPS |  | 10 SPS | SINC3 | 0.176 | 0.088 | 0.043 | 0.028 | 0.018 | 0.014 | 14400 | - | - | - | - | 4317.77 | -
19200 SPS |  | 10 SPS | SINC4 | 0.164 | 0.076 | 0.04 | 0.024 | 0.016 | 0.013 | 19200 | - | - | - | - | 5757.53 | -
38400 SPS |  | 16.6 SPS | SINC1 | 0.306 | 0.147 | 0.077 | 0.044 | 0.03 | 0.024 | 38400 | - | - | - | - | 11521.75 | -
|  | 16.6 SPS | SINC2 | 0.248 | 0.122 | 0.068 | 0.037 | 0.024 | 0.02 | *Integrated frequency response up to 120kHz with 0.1Hz resolution
|  | 16.6 SPS | SINC3 | 0.216 | 0.12 | 0.06 | 0.033 | 0.022 | 0.017
|  | 16.6 SPS | SINC4 | 0.214 | 0.101 | 0.054 | 0.031 | 0.022 | 0.016
|  | 20 SPS | FIR | 0.393 | 0.191 | 0.104 | 0.057 | 0.039 | 0.03
|  | 20 SPS | SINC1 | 0.336 | 0.167 | 0.085 | 0.049 | 0.033 | 0.026
|  | 20 SPS | SINC2 | 0.27 | 0.136 | 0.07 | 0.039 | 0.028 | 0.021
|  | 20 SPS | SINC3 | 0.237 | 0.124 | 0.067 | 0.035 | 0.024 | 0.02
|  | 20 SPS | SINC4 | 0.229 | 0.113 | 0.06 | 0.034 | 0.022 | 0.017
|  | 50 SPS | SINC1 | 0.514 | 0.255 | 0.14 | 0.077 | 0.051 | 0.042
|  | 50 SPS | SINC2 | 0.426 | 0.209 | 0.108 | 0.064 | 0.042 | 0.033
|  | 50 SPS | SINC3 | 0.389 | 0.196 | 0.104 | 0.057 | 0.038 | 0.03
|  | 50 SPS | SINC4 | 0.358 | 0.175 | 0.096 | 0.055 | 0.036 | 0.028
|  | 60 SPS | SINC1 | 0.558 | 0.285 | 0.151 | 0.085 | 0.055 | 0.045
|  | 60 SPS | SINC2 | 0.465 | 0.235 | 0.121 | 0.068 | 0.046 | 0.036
|  | 60 SPS | SINC3 | 0.414 | 0.208 | 0.112 | 0.064 | 0.042 | 0.034
|  | 60 SPS | SINC4 | 0.383 | 0.195 | 0.105 | 0.059 | 0.04 | 0.031
|  | 100 SPS | SINC1 | 0.734 | 0.361 | 0.192 | 0.108 | 0.071 | 0.058
|  | 100 SPS | SINC2 | 0.604 | 0.305 | 0.156 | 0.088 | 0.059 | 0.048
|  | 100 SPS | SINC3 | 0.531 | 0.277 | 0.143 | 0.081 | 0.054 | 0.043
|  | 100 SPS | SINC4 | 0.511 | 0.255 | 0.134 | 0.076 | 0.05 | 0.041
|  | 400 SPS | SINC1 | 1.438 | 0.734 | 0.38 | 0.215 | 0.143 | 0.116
|  | 400 SPS | SINC2 | 1.186 | 0.607 | 0.313 | 0.178 | 0.119 | 0.095
|  | 400 SPS | SINC3 | 1.072 | 0.55 | 0.285 | 0.161 | 0.107 | 0.087
|  | 400 SPS | SINC4 | 0.995 | 0.508 | 0.266 | 0.151 | 0.101 | 0.081
|  | 1200 SPS | SINC1 | 2.451 | 1.254 | 0.651 | 0.368 | 0.244 | 0.197
|  | 1200 SPS | SINC2 | 2.038 | 1.037 | 0.545 | 0.309 | 0.205 | 0.165
|  | 1200 SPS | SINC3 | 1.858 | 0.96 | 0.494 | 0.281 | 0.186 | 0.148
|  | 1200 SPS | SINC4 | 1.743 | 0.89 | 0.459 | 0.261 | 0.174 | 0.139
|  | 2400 SPS | SINC1 | 3.411 | 1.724 | 0.903 | 0.51 | 0.335 | 0.27
|  | 2400 SPS | SINC2 | 2.87 | 1.468 | 0.77 | 0.435 | 0.286 | 0.23
|  | 2400 SPS | SINC3 | 2.656 | 1.337 | 0.705 | 0.395 | 0.262 | 0.211
|  | 2400 SPS | SINC4 | 2.475 | 1.262 | 0.657 | 0.371 | 0.245 | 0.198
|  | 4800 SPS | SINC1 | 4.59 | 2.329 | 1.221 | 0.682 | 0.446 | 0.361
|  | 4800 SPS | SINC2 | 4.091 | 2.07 | 1.077 | 0.606 | 0.398 | 0.321
|  | 4800 SPS | SINC3 | 3.72 | 1.894 | 0.998 | 0.56 | 0.367 | 0.297
|  | 4800 SPS | SINC4 | 3.535 | 1.784 | 0.926 | 0.527 | 0.349 | 0.277
|  | 7200 SPS | SINC1 | 5.326 | 2.709 | 1.407 | 0.792 | 0.516 | 0.409
|  | 7200 SPS | SINC2 | 4.867 | 2.467 | 1.28 | 0.726 | 0.472 | 0.379
|  | 7200 SPS | SINC3 | 4.567 | 2.31 | 1.209 | 0.682 | 0.445 | 0.359
|  | 7200 SPS | SINC4 | 4.365 | 2.211 | 1.143 | 0.642 | 0.426 | 0.341
|  | 14400 SPS | SINC5 | 6.377 | 3.235 | 1.675 | 0.929 | 0.596 | 0.466
|  | 19200 SPS | SINC5 | 8.72 | 4.432 | 2.285 | 1.227 | 0.747 | 0.555
|  | 38400 SPS | SINC5 | 103.55 | 51.76 | 25.95 | 13.02 | 6.493 | 3.276