## Summary of Results
Initial Reference Voltages
VREF_INIT_ADC: | 5.0484 | [V] |  |  | NOTE: The large offset and gain coefficents mainly correct for the bridge's offset and gain error. The ADC offset and gain errors are included in these values; however, the bridge errors dominate.
VREF_INIT_DMM: | 5.0736 | [V]
Calibration Coefficients
OFFSETDC: | 0.00011133 | [V] | GAINDC: |  | 0.99707 |  | [V]
OFFSETAC: | 0.00011219 | [V] | GAINAC: |  | 0.9971 |  | [V]
OFFSETDMM: | 0.00011266 | [V] | GAINDMM: |  | 0.99229 |  | [V]
Error vs. temperature with 0 V Input (SW1 & SW4 Open)
Temperature | VOS_DC_<X>°C | VOS_AC_<X>°C |  | VOS_DMM_<X>ᵒC |  | VOS_DC_<X>°C - VOS_DMM_<X>ᵒC |  |  | VOS_AC_<X>°C - VOS_DMM_<X>ᵒC
-40 | -0.0003059 | -0.0003067 |  | -0.00030691 |  | 1004.6239 |  |  | 207.2652
-20 | -0.00022401 | -0.00022467 |  | -0.00022506 |  | 1052.8254 |  |  | 394.067
0 | -0.00013787 | -0.00013731 |  | -0.00013779 |  | -80.1382 |  |  | 484.5112
25 | 0 | 0 |  | 0 |  | 0 |  |  | 0
40 | -1.2537e-05 | -1.3261e-05 |  | -1.3507e-05 |  | 970.0496 |  |  | 245.8215
80 | 0.00010673 | 0.00010288 |  | 0.00010338 |  | 3351.9199 |  |  | -497.731
100 | 0.00014994 | 0.00014486 |  | 0.00014547 |  | 4471.958 |  |  | -608.4218
125 | 0.00020642 | 0.00019846 |  | 0.00019924 |  | 7178.5558 |  |  | -785.21
[5 empty] | Error vs. Temperature Slope: | 43.9921 |  |  | 7.6953
Error vs. temperature with 10 mV Input (SW1 & SW4 Closed)
Temperature | VFS_DC_<X>°C | VFS_AC_<X>°C |  | VFS_DMM_<X>ᵒC |  | VOS_DC_<X>°C - VOS_DMM_<X>ᵒC |  |  | VOS_AC_<X>°C - VOS_DMM_<X>ᵒC
-40 | 0.0097549 | 0.0097562 |  | 0.0097569 |  | -1993.8228 |  |  | -644.1125
-20 | 0.0098447 | 0.0098455 |  | 0.0098458 |  | -1097.2765 |  |  | -269.2133
0 | 0.0099207 | 0.0099216 |  | 0.0099218 |  | -1107.0719 |  |  | -189.4872
25 | 0.01 | 0.01 |  | 0.01 |  | 0 |  |  | 0
40 | 0.010045 | 0.010044 |  | 0.010044 |  | 1084.8401 |  |  | 95.1966
80 | 0.010155 | 0.010152 |  | 0.010152 |  | 3464.316 |  |  | 380.4365
100 | 0.010202 | 0.010197 |  | 0.010197 |  | 4737.8544 |  |  | 388.4148
125 | 0.010263 | 0.010257 |  | 0.010256 |  | 7268.1413 |  |  | 868.7848
[5 empty] | Error vs. Temperature Slope: | 56.1331 |  |  | 9.1691
Noise vs. Temp (0 V Input) |  |  |  |  | Full-Scale Load Cell Input: |  |  | 0.01
Temperature | VOS_DC_<X>°C | VOS_AC_<X>°C (Averaged) |  |  | Noise-Free Counts_DC |  | Noise-Free Counts_AC
-40 | 202.6857 | 172.7861 |  |  | 49337.4705 |  | 57875.0345
-20 | 239.8582 | 136.238 |  |  | 41691.2959 |  | 73400.9605
0 | 246.6169 | 190.3438 |  |  | 40548.7284 |  | 52536.5038
25 | 114.0153 | 115.2274 |  |  | 87707.5266 |  | 86784.8781
40 | 231.7772 | 153.0611 |  |  | 43144.8752 |  | 65333.3723
80 | 282.8343 | 231.1895 |  |  | 35356.3847 |  | 43254.5539
100 | 552.4453 | 358.7588 |  |  | 18101.3406 |  | 27873.8776
125 | 752.2659 | 239.895 |  |  | 13293.172 |  | 41684.9123
Noise vs. Temp (10 mV Input) |  |  |  |  | Full-Scale Load Cell Input: |  |  | 0.01
Temperature | VFS_DC_-40°C | VFS_AC_-40°C (Averaged) |  |  | Noise-Free Counts_DC |  | Noise-Free Counts_AC
-40 | 258.8853 | 123.2717 |  |  | 38627.1513 |  | 81121.6217
-20 | 323.4596 | 106.9628 |  |  | 30915.7577 |  | 93490.4404
0 | 269.9048 | 147.7718 |  |  | 37050.1038 |  | 67671.927
25 | 184.3933 | 93.0047 |  |  | 54231.9049 |  | 107521.3912
40 | 294.0008 | 162.9052 |  |  | 34013.5136 |  | 61385.3805
80 | 472.2966 | 286.1769 |  |  | 21173.1344 |  | 34943.4171
100 | 701.6496 | 520.1213 |  |  | 14252.1287 |  | 19226.2827
125 | 632.0062 | 628.8105 |  |  | 15822.6295 |  | 15903.0412
NOTE: The ac-excited bridge noise is calculated as an average of samples taken between positive and negative excitation voltage polarities. Without this averaging, noise performance is comparable to the dc-excited noise performance.
↓ 50,000 counts ↓ noise-free limit
↓ 50,000 counts ↓ noise-free limit

## TIPD188 Data
Offset-calibration (SW1 & SW4 Open)
Measure the initial ADC’s reference voltage (using the ADC): | [12 empty] | VREF_INIT_ADC: |  | 5.0484 |  | [V]
ADC Reference (+ Polarity): |  | 5.0484 | [V] |  | (The ADC measurements are ratiometric; therefore, only one reference polarity measurement is required)
Measure the initial ADC’s reference voltage (using the DMM): | [12 empty] | VREF_INIT_DMM: |  | 5.0736 |  | [V]
DMM2 Reference (+ Polarity): |  | 5.0736 | [V] |  | (DMM measurements will be compared and scaled to this reference measurement, to simulate ratiometric measurements)
Measure the (average) dc excited bridge output offset voltage with the ADC: | [12 empty] | OFFSETDC: |  | 0.00011133 |  | [V]
ADC Codes (+ Polarity): |  | 00171DA6 | 00172210 | 001721EE | 00171CB7 | 00171E5F | 001722C7 | 00171E54 | 00171E31 | 00172114 | 00172010 | 00171ED5 | 00171CB9 | 0017201E | 00171FDC | 00171E46 | 001720D0
Voltages (+ Polarity): |  | 0.00011129 | 0.00011137 | 0.00011137 | 0.00011127 | 0.0001113 | 0.00011139 | 0.0001113 | 0.0001113 | 0.00011136 | 0.00011134 | 0.00011131 | 0.00011127 | 0.00011134 | 0.00011133 | 0.0001113 | 0.00011135
Measure the (average) ac excited bridge output offset voltages with the ADC: | [12 empty] | OFFSETAC: |  | 0.00011219 |  | [V]
ADC Codes (+ Polarity): |  | 00172236 | 00172294 | 0017261B | 001720C9 | 00171F55 | 00172034 | 00171EE3 | 001721A1 | 0017224F | 00172389 | 0017232D | 001723CF | 0017250F | 00172097 | 00172393 | 00172364
Voltages (+ Polarity): |  | 0.00011138 | 0.00011138 | 0.00011145 | 0.00011135 | 0.00011132 | 0.00011134 | 0.00011131 | 0.00011137 | 0.00011138 | 0.0001114 | 0.0001114 | 0.00011141 | 0.00011143 | 0.00011135 | 0.0001114 | 0.0001114
ADC Codes (- Polarity): |  | FFE88896 | FFE88C31 | FFE88933 | FFE8877A | FFE883FE | FFE884DA | FFE8896B | FFE8861D | FFE88813 | FFE8854E | FFE882CF | FFE88633 | FFE887E8 | FFE88999 | FFE88C01 | FFE88939
Voltages (- Polarity): |  | -0.00011298 | -0.00011291 | -0.00011297 | -0.000113 | -0.00011307 | -0.00011305 | -0.00011296 | -0.00011303 | -0.00011299 | -0.00011304 | -0.00011309 | -0.00011302 | -0.00011299 | -0.00011296 | -0.00011292 | -0.00011297
Measure the (average) ac excited bridge output offset voltage with DMM1: | [12 empty] | OFFSETDMM: |  | 0.00011266 |  | [V]
VOS_DMM_P: |  | 0.00042051 | [V] |  | VREF_OS_DMM_P: |  | 5.0736 | [V] | → (Eq. 5) → |  | VOS_DMM_P_NORM: |  | 0.00042052 | [V] | ↑ (Eq. 7) ↑
VOS_DMM_N: |  | 0.00019525 | [V] |  | VREF_OS_DMM_N: |  | -5.0749 | [V] | → (Eq. 6) → |  | VOS_DMM_P_NORM: |  | 0.0001952 | [V]
Bridge measurment error vs. temperature with 0 V Input (SW1 & SW4 Open)
Measure the (average) dc excited bridge output offset voltage with the ADC, for all desired temperature settings:
ADC Codes (+ Polarity, -40ᵒC): |  | FFD791AB | FFD78ED9 | FFD79350 | FFD79545 | FFD79721 | FFD79913 | FFD7950F | FFD798A0 | FFD79796 | FFD796B9 | FFD79236 | FFD79979 | FFD79678 | FFD79818 | FFD79452 | FFD799A0
Voltages (+ Polarity, -40ᵒC): |  | -0.00019466 | -0.00019471 | -0.00019462 | -0.00019459 | -0.00019455 | -0.00019452 | -0.00019459 | -0.00019452 | -0.00019454 | -0.00019456 | -0.00019464 | -0.00019451 | -0.00019456 | -0.00019453 | -0.00019461 | -0.00019451
Average (+ Polarity, -40ᵒC): |  | -0.00019458 |  | [V] | → (Subtract OFFSETDC) → |  |  | -0.0003059 |  | [V] |  |  | VOS_DC_-40°C |  | -0.0003059 |  | [V]
ADC Codes (+ Polarity, -20ᵒC): |  | FFE89295 | FFE8982A | FFE8992B | FFE89439 | FFE89A2C | FFE89BE3 | FFE89C3A | FFE89C61 | FFE89BBA | FFE89645 | FFE8937B | FFE892FD | FFE899AD | FFE89F56 | FFE89704 | FFE89341
Voltages (+ Polarity, -20ᵒC): |  | -0.00011279 | -0.00011269 | -0.00011267 | -0.00011276 | -0.00011265 | -0.00011262 | -0.00011261 | -0.00011261 | -0.00011262 | -0.00011272 | -0.00011277 | -0.00011278 | -0.00011266 | -0.00011255 | -0.00011271 | -0.00011278
Average (+ Polarity, -20ᵒC): |  | -0.00011269 |  | [V] | → (Subtract OFFSETDC) → |  |  | -0.00022401 |  | [V] |  |  | VOS_DC_-20°C |  | -0.00022401 |  | [V]
ADC Codes (+ Polarity, 0ᵒC): |  | FFFA7F3A | FFFA7EBF | FFFA75D8 | FFFA7EFE | FFFA79CF | FFFA7A50 | FFFA82F5 | FFFA780D | FFFA7914 | FFFA7FDE | FFFA76FE | FFFA7C58 | FFFA7AEC | FFFA826C | FFFA7B9E | FFFA7F7E
Voltages (+ Polarity, 0ᵒC): |  | -2.6494e-05 | -2.6503e-05 | -2.6671e-05 | -2.6499e-05 | -2.6596e-05 | -2.6587e-05 | -2.6424e-05 | -2.6629e-05 | -2.661e-05 | -2.6482e-05 | -2.6649e-05 | -2.6549e-05 | -2.6575e-05 | -2.6434e-05 | -2.6562e-05 | -2.6489e-05
Average (+ Polarity, 0ᵒC): |  | -2.6547e-05 |  | [V] | → (Subtract OFFSETDC) → |  |  | -0.00013787 |  | [V] |  |  | VOS_DC_0°C |  | -0.00013787 |  | [V]
ADC Codes (+ Polarity, 40ᵒC): |  | 00147F72 | 001483EB | 00148259 | 00148389 | 00148668 | 0014822E | 001484C1 | 00148400 | 00148A82 | 00148962 | 001486EB | 00148521 | 001489DB | 00148861 | 001482EF | 00147E2F
Voltages (+ Polarity, 40ᵒC): |  | 9.8687e-05 | 9.8771e-05 | 9.8741e-05 | 9.8764e-05 | 9.8818e-05 | 9.8738e-05 | 9.8787e-05 | 9.8773e-05 | 9.8895e-05 | 9.8874e-05 | 9.8827e-05 | 9.8794e-05 | 9.8883e-05 | 9.8855e-05 | 9.8752e-05 | 9.8663e-05
Average (+ Polarity, 40ᵒC): |  | 9.8789e-05 |  | [V] | → (Subtract OFFSETDC) → |  |  | -1.2537e-05 |  | [V] |  |  | VOS_DC_40°C |  | -1.2537e-05 |  | [V]
ADC Codes (+ Polarity, 80ᵒC): |  | 002D534F | 002D4CB0 | 002D4BD0 | 002D4EF7 | 002D47F1 | 002D466C | 002D4D3B | 002D4445 | 002D4700 | 002D4D13 | 002D4A8B | 002D4EF8 | 002D464E | 002D4DE0 | 002D4695 | 002D494D
Voltages (+ Polarity, 80ᵒC): |  | 0.00021822 | 0.00021809 | 0.00021808 | 0.00021814 | 0.00021801 | 0.00021798 | 0.00021811 | 0.00021794 | 0.00021799 | 0.0002181 | 0.00021805 | 0.00021814 | 0.00021797 | 0.00021812 | 0.00021798 | 0.00021803
Average (+ Polarity, 80ᵒC): |  | 0.00021806 |  | [V] | → (Subtract OFFSETDC) → |  |  | 0.00010673 |  | [V] |  |  | VOS_DC_80°C |  | 0.00010673 |  | [V]
ADC Codes (+ Polarity, 100ᵒC): |  | 0036488D | 003649A2 | 003644A2 | 00364EE2 | 00363182 | 0036370D | 003646D3 | 00364DC8 | 00364088 | 0036420B | 003645C4 | 00364393 | 003643EF | 00364AA8 | 00363E45 | 00364A96
Voltages (+ Polarity, 100ᵒC): |  | 0.00026135 | 0.00026137 | 0.00026127 | 0.00026147 | 0.00026091 | 0.00026102 | 0.00026132 | 0.00026145 | 0.0002612 | 0.00026123 | 0.0002613 | 0.00026125 | 0.00026126 | 0.00026139 | 0.00026115 | 0.00026139
Average (+ Polarity, 100ᵒC): |  | 0.00026127 |  | [V] | → (Subtract OFFSETDC) → |  |  | 0.00014994 |  | [V] |  |  | VOS_DC_100°C |  | 0.00014994 |  | [V]
ADC Codes (+ Polarity, 125ᵒC): |  | 00420211 | 0041EF58 | 00420496 | 0041F6D9 | 0041FB2B | 0041F031 | 004206FD | 004215BF | 004200B3 | 0041FFD3 | 0042028F | 0042003A | 00420AED | 004204AD | 004200B9 | 0041EDBF
Voltages (+ Polarity, 125ᵒC): |  | 0.0003178 | 0.00031744 | 0.00031784 | 0.00031758 | 0.00031767 | 0.00031746 | 0.00031789 | 0.00031817 | 0.00031777 | 0.00031775 | 0.00031781 | 0.00031776 | 0.00031796 | 0.00031785 | 0.00031777 | 0.00031741
Average (+ Polarity, 125ᵒC): |  | 0.00031775 |  | [V] | → (Subtract OFFSETDC) → |  |  | 0.00020642 |  | [V] |  |  | VOS_DC_125°C |  | 0.00020642 |  | [V]
Measure the (average) ac excited bridge output offset voltage with the ADC, for all desired temperature settings:
ADC Codes (+ Polarity, -40ᵒC): |  | FFD771A4 | FFD7722F | FFD771CA | FFD76E25 | FFD77638 | FFD77039 | FFD76F8F | FFD76CDA | FFD76EA9 | FFD773D4 | FFD76EFB | FFD770B4 | FFD769C5 | FFD7700D | FFD772DA | FFD770A6
Voltages (+ Polarity, -40ᵒC): |  | -0.00019526 | -0.00019525 | -0.00019525 | -0.00019532 | -0.00019517 | -0.00019528 | -0.0001953 | -0.00019535 | -0.00019531 | -0.00019522 | -0.00019531 | -0.00019528 | -0.00019541 | -0.00019529 | -0.00019523 | -0.00019528
| ADC Codes (- Polarity, -40ᵒC): | 00283AE9 | 00283803 | 0028381C | 00283936 | 00283F2D | 0028391D | 0028403D | 00283D72 | 00283C96 | 0028391C | 00283C3E | 00284741 | 00284816 | 00283D02 | 002846E1 | 00283DD1
Voltages (- Polarity, -40ᵒC): |  | 0.00019369 | 0.00019363 | 0.00019364 | 0.00019366 | 0.00019377 | 0.00019365 | 0.00019379 | 0.00019374 | 0.00019372 | 0.00019365 | 0.00019371 | 0.00019392 | 0.00019394 | 0.00019373 | 0.00019391 | 0.00019374
Average (+ Polarity, -40ᵒC): |  | -0.00019528 |  | Average (- Polarity, -40ᵒC): |  |  | 0.00019374 |  | → (Eq. 2) & (Subtract OFFSETDC) → |  |  |  | VOS_AC_-40°C |  | -0.0003067 |  | [V]
ADC Codes (+ Polarity, -20ᵒC): |  | FFE84373 | FFE83F04 | FFE8445A | FFE84506 | FFE843A4 | FFE84299 | FFE84261 | FFE83D5E | FFE840DC | FFE8421F | FFE846AC | FFE842DF | FFE844A0 | FFE843CC | FFE8414B | FFE83E37
Voltages (+ Polarity, -20ᵒC): |  | -0.00011428 | -0.00011436 | -0.00011426 | -0.00011425 | -0.00011428 | -0.0001143 | -0.0001143 | -0.00011439 | -0.00011433 | -0.0001143 | -0.00011422 | -0.00011429 | -0.00011426 | -0.00011427 | -0.00011432 | -0.00011438
| ADC Codes (- Polarity, -20ᵒC): | 0016FF09 | 0016FBE0 | 0016FD72 | 0016F6EB | 0017011E | 0016FC53 | 0016F75D | 001700D8 | 0016FB57 | 0016F53B | 0016FEDE | 001702F8 | 0016FC8F | 0016FAD6 | 0016FE22 | 0016FA8F
Voltages (- Polarity, -20ᵒC): |  | 0.00011072 | 0.00011066 | 0.00011069 | 0.00011056 | 0.00011075 | 0.00011066 | 0.00011057 | 0.00011075 | 0.00011065 | 0.00011053 | 0.00011071 | 0.00011079 | 0.00011067 | 0.00011064 | 0.0001107 | 0.00011063
Average (+ Polarity, -20ᵒC): |  | -0.0001143 |  | Average (- Polarity, -20ᵒC): |  |  | 0.00011067 |  | → (Eq. 2) & (Subtract OFFSETDC) → |  |  |  | VOS_AC_-20°C |  | -0.00022467 |  | [V]
ADC Codes (+ Polarity, 0ᵒC): |  | FFFA7966 | FFFA7DB6 | FFFA7DE9 | FFFA7D13 | FFFA77C3 | FFFA7F8E | FFFA7F4F | FFFA7B7F | FFFA7E4C | FFFA7E13 | FFFA7EBB | FFFA75CE | FFFA8316 | FFFA7E9A | FFFA7EF2 | FFFA7FC5
Voltages (+ Polarity, 0ᵒC): |  | -2.6604e-05 | -2.6523e-05 | -2.6519e-05 | -2.6535e-05 | -2.6635e-05 | -2.6488e-05 | -2.6493e-05 | -2.6564e-05 | -2.6512e-05 | -2.6516e-05 | -2.6504e-05 | -2.6671e-05 | -2.6422e-05 | -2.6506e-05 | -2.65e-05 | -2.6484e-05
| ADC Codes (- Polarity, 0ᵒC): | 0004E722 | 0004ED8F | 0004F05D | 0004F1EB | 0004E766 | 0004F3D9 | 0004F13A | 0004E9FA | 0004EBB9 | 0004EB7F | 0004EDF7 | 0004EE5C | 0004E37D | 0004F00C | 0004E8DC | 0004EF64
Voltages (- Polarity, 0ᵒC): |  | 2.3605e-05 | 2.3726e-05 | 2.3778e-05 | 2.3808e-05 | 2.361e-05 | 2.3844e-05 | 2.3795e-05 | 2.3658e-05 | 2.3691e-05 | 2.3687e-05 | 2.3733e-05 | 2.3741e-05 | 2.3536e-05 | 2.3772e-05 | 2.3637e-05 | 2.376e-05
Average (+ Polarity, 0ᵒC): |  | -2.653e-05 |  | Average (- Polarity, 0ᵒC): |  |  | 2.3711e-05 |  | → (Eq. 2) & (Subtract OFFSETDC) → |  |  |  | VOS_AC_-0°C |  | -0.00013731 |  | [V]
ADC Codes (+ Polarity, 40ᵒC): |  | 0014A3C4 | 0014A8C9 | 0014A2A4 | 0014A2FA | 0014AAFF | 0014AA08 | 0014A5A8 | 0014A41E | 0014A7F3 | 0014A1C2 | 0014A61B | 0014A644 | 0014A121 | 0014A82A | 0014A3D3 | 0014ACC8
Voltages (+ Polarity, 40ᵒC): |  | 9.937e-05 | 9.9464e-05 | 9.9349e-05 | 9.9355e-05 | 9.9506e-05 | 9.9488e-05 | 9.9405e-05 | 9.9377e-05 | 9.9449e-05 | 9.9332e-05 | 9.9414e-05 | 9.9417e-05 | 9.932e-05 | 9.9453e-05 | 9.9371e-05 | 9.9539e-05
| ADC Codes (- Polarity, 40ᵒC): | FFEB8F59 | FFEB8FD9 | FFEB8F57 | FFEB8D67 | FFEB90EC | FFEB8F34 | FFEB8C2C | FFEB8B2E | FFEB8B74 | FFEB8DAD | FFEB8E43 | FFEB8DE0 | FFEB8A83 | FFEB8E3B | FFEB8B4C | FFEB8E1B
Voltages (- Polarity, 40ᵒC): |  | -9.8409e-05 | -9.8399e-05 | -9.8409e-05 | -9.8445e-05 | -9.8379e-05 | -9.8411e-05 | -9.8468e-05 | -9.8487e-05 | -9.8482e-05 | -9.844e-05 | -9.8429e-05 | -9.8436e-05 | -9.85e-05 | -9.843e-05 | -9.8485e-05 | -9.8432e-05
Average (+ Polarity, 40ᵒC): |  | 9.9413e-05 |  | Average (- Polarity, 40ᵒC): |  |  | -9.844e-05 |  | → (Eq. 2) & (Subtract OFFSETDC) → |  |  |  | VOS_AC_40°C |  | -1.3261e-05 |  | [V]
ADC Codes (+ Polarity, 80ᵒC): |  | 002D5306 | 002D5FBF | 002D520D | 002D5A0B | 002D5B6F | 002D5339 | 002D52FF | 002D5F5F | 002D4FE6 | 002D5133 | 002D504A | 002D50D4 | 002D472C | 002D481C | 002D5401 | 002D47BA
Voltages (+ Polarity, 80ᵒC): |  | 0.00021821 | 0.00021845 | 0.0002182 | 0.00021835 | 0.00021837 | 0.00021822 | 0.00021821 | 0.00021845 | 0.00021816 | 0.00021818 | 0.00021816 | 0.00021817 | 0.00021799 | 0.00021801 | 0.00021823 | 0.000218
| ADC Codes (- Polarity, 80ᵒC): | FFD3F65E | FFD3F668 | FFD3F993 | FFD3EF0C | FFD3FCEB | FFD3FBD9 | FFD3FECE | FFD3FD15 | FFD407C6 | FFD3F7D8 | FFD3FADD | FFD3FF3E | FFD3FCB2 | FFD3FF7E | FFD3F29E | FFD3FD96
Voltages (- Polarity, 80ᵒC): |  | -0.00021202 | -0.00021202 | -0.00021196 | -0.00021216 | -0.0002119 | -0.00021192 | -0.00021186 | -0.00021189 | -0.00021169 | -0.00021199 | -0.00021193 | -0.00021185 | -0.0002119 | -0.00021185 | -0.00021209 | -0.00021188
Average (+ Polarity, 80ᵒC): |  | 0.00021821 |  | Average (- Polarity, 80ᵒC): |  |  | -0.00021193 |  | → (Eq. 2) & (Subtract OFFSETDC) → |  |  |  | VOS_AC_80°C |  | 0.00010288 |  | [V]
ADC Codes (+ Polarity, 100ᵒC): |  | 00365276 | 00365161 | 00364637 | 003649A0 | 0036453C | 00364C6A | 003647AA | 00364195 | 00364675 | 00364A5B | 003648D4 | 003649D5 | 003636E8 | 0036488D | 00364125 | 0036496F
Voltages (+ Polarity, 100ᵒC): |  | 0.00026153 | 0.00026151 | 0.0002613 | 0.00026137 | 0.00026129 | 0.00026142 | 0.00026133 | 0.00026122 | 0.00026131 | 0.00026138 | 0.00026135 | 0.00026137 | 0.00026102 | 0.00026135 | 0.00026121 | 0.00026136
| ADC Codes (- Polarity, 100ᵒC): | FFCB8290 | FFCB78C3 | FFCB8339 | FFCB867B | FFCB8316 | FFCB7C0E | FFCB817F | FFCB7F73 | FFCB81B8 | FFCB85A6 | FFCB7B5F | FFCB7526 | FFCB77F7 | FFCB7E0C | FFCB7F5C | FFCB8700
Voltages (- Polarity, 100ᵒC): |  | -0.00025271 | -0.0002529 | -0.0002527 | -0.00025264 | -0.0002527 | -0.00025284 | -0.00025273 | -0.00025277 | -0.00025273 | -0.00025266 | -0.00025285 | -0.00025297 | -0.00025291 | -0.0002528 | -0.00025277 | -0.00025263
Average (+ Polarity, 100ᵒC): |  | 0.00026133 |  | Average (- Polarity, 100ᵒC): |  |  | -0.00025277 |  | → (Eq. 2) & (Subtract OFFSETDC) → |  |  |  | VOS_AC_100°C |  | 0.00014486 |  | [V]
ADC Codes (+ Polarity, 125ᵒC): |  | 0041F371 | 0041F05B | 0041EB36 | 0041F07F | 0041EFB3 | 0041F355 | 0041F82E | 0041F568 | 0041F07E | 0041F353 | 0041F370 | 0041F81E | 0041EF22 | 0041FB84 | 0041F7E2 | 0041E876
Voltages (+ Polarity, 125ᵒC): |  | 0.00031752 | 0.00031746 | 0.00031737 | 0.00031747 | 0.00031745 | 0.00031752 | 0.00031761 | 0.00031756 | 0.00031747 | 0.00031752 | 0.00031752 | 0.00031761 | 0.00031744 | 0.00031767 | 0.0003176 | 0.00031731
| ADC Codes (- Polarity, 125ᵒC): | FFC0DCDD | FFC0E2A9 | FFC0E47D | FFC0E43A | FFC0DCC8 | FFC0F2A9 | FFC0EB70 | FFC0EAE0 | FFC0EB9C | FFC0E9A4 | FFC0EE57 | FFC0E429 | FFC0E750 | FFC0E102 | FFC0EA5F | FFC0E94F
Voltages (- Polarity, 125ᵒC): |  | -0.00030397 | -0.00030387 | -0.00030383 | -0.00030384 | -0.00030398 | -0.00030356 | -0.0003037 | -0.00030371 | -0.0003037 | -0.00030373 | -0.00030365 | -0.00030384 | -0.00030378 | -0.0003039 | -0.00030372 | -0.00030374
Average (+ Polarity, 125ᵒC): |  | 0.00031751 |  | Average (- Polarity, 125ᵒC): |  |  | -0.00030378 |  | → (Eq. 2) & (Subtract OFFSETDC) → |  |  |  | VOS_AC_125°C |  | 0.00019846 |  | [V]
Measure the (average) ac excited bridge output offset voltages with DMM1, for all desired temperature settings:
VOS_DMM_-40ᵒC_P: |  | -0.00019384 | [V] | VREF_OS_DMM_-40ᵒC_P: |  | 5.079 | [V] | →(Eq. 8)→ | VOS_DMM_-40ᵒC_P_NORM: |  | -0.00019364 | [V] |  |  | ↓ (Eq. 10) ↓
VOS_DMM_-40ᵒC_N: |  | 0.0001951 | [V] | VREF_OS_DMM_-40ᵒC_N: |  | -5.08 | [V] | →(Eq. 9)→ | VOS_DMM_-40ᵒC_P_NORM: |  | 0.00019486 | [V] | VOS_DMM_-40ᵒC: |  | -0.00030691 |  | [V]
VOS_DMM_-20ᵒC_P: |  | -0.00011141 | [V] | VREF_OS_DMM_-20ᵒC_P: |  | 5.0774 | [V] | →(Eq. 8)→ | VOS_DMM_-20ᵒC_P_NORM: |  | -0.00011133 | [V] |  |  | ↓ (Eq. 10) ↓
VOS_DMM_-20ᵒC_N: |  | 0.00011359 | [V] | VREF_OS_DMM_-20ᵒC_N: |  | -5.0785 | [V] | →(Eq. 9)→ | VOS_DMM_-20ᵒC_P_NORM: |  | 0.00011348 | [V] | VOS_DMM_-20ᵒC: |  | -0.00022506 |  | [V]
VOS_DMM_0ᵒC_P: |  | -2.35e-05 | [V] | VREF_OS_DMM_0ᵒC_P: |  | 5.0756 | [V] | →(Eq. 8)→ | VOS_DMM_0ᵒC_P_NORM: |  | -2.3491e-05 | [V] |  |  | ↓ (Eq. 10) ↓
VOS_DMM_0ᵒC_N: |  | 2.679e-05 | [V] | VREF_OS_DMM_0ᵒC_N: |  | -5.0768 | [V] | →(Eq. 9)→ | VOS_DMM_0ᵒC_P_NORM: |  | 2.6773e-05 | [V] | VOS_DMM_0ᵒC: |  | -0.00013779 |  | [V]
VOS_DMM_40ᵒC_P: |  | 9.751e-05 | [V] | VREF_OS_DMM_40ᵒC_P: |  | 5.0728 | [V] | →(Eq. 8)→ | VOS_DMM_40ᵒC_P_NORM: |  | 9.7526e-05 | [V] |  |  | ↓ (Eq. 10) ↓
VOS_DMM_40ᵒC_N: |  | -0.00010079 | [V] | VREF_OS_DMM_40ᵒC_N: |  | -5.0741 | [V] | →(Eq. 9)→ | VOS_DMM_40ᵒC_P_NORM: |  | -0.00010078 | [V] | VOS_DMM_40ᵒC: |  | -1.3507e-05 |  | [V]
VOS_DMM_80ᵒC_P: |  | 0.00021231 | [V] | VREF_OS_DMM_80ᵒC_P: |  | 5.069 | [V] | →(Eq. 8)→ | VOS_DMM_80ᵒC_P_NORM: |  | 0.0002125 | [V] |  |  | ↓ (Eq. 10) ↓
VOS_DMM_80ᵒC_N: |  | -0.00021944 | [V] | VREF_OS_DMM_80ᵒC_N: |  | -5.0704 | [V] | →(Eq. 9)→ | VOS_DMM_80ᵒC_P_NORM: |  | -0.00021958 | [V] | VOS_DMM_80ᵒC: |  | 0.00010338 |  | [V]
VOS_DMM_100ᵒC_P: |  | 0.00025639 | [V] | VREF_OS_DMM_100ᵒC_P: |  | 5.0671 | [V] | →(Eq. 8)→ | VOS_DMM_100ᵒC_P_NORM: |  | 0.00025672 | [V] |  |  | ↓ (Eq. 10) ↓
VOS_DMM_100ᵒC_N: |  | -0.00025929 | [V] | VREF_OS_DMM_100ᵒC_N: |  | -5.0687 | [V] | →(Eq. 9)→ | VOS_DMM_100ᵒC_P_NORM: |  | -0.00025954 | [V] | VOS_DMM_100ᵒC: |  | 0.00014547 |  | [V]
VOS_DMM_125ᵒC_P: |  | 0.00031177 | [V] | VREF_OS_DMM_125ᵒC_P: |  | 5.0647 | [V] | →(Eq. 8)→ | VOS_DMM_125ᵒC_P_NORM: |  | 0.00031232 | [V] |  |  | ↓ (Eq. 10) ↓
VOS_DMM_125ᵒC_N: |  | -0.00031104 | [V] | VREF_OS_DMM_125ᵒC_N: |  | -5.0664 | [V] | →(Eq. 9)→ | VOS_DMM_125ᵒC_P_NORM: |  | -0.00031149 | [V] | VOS_DMM_125ᵒC: |  | 0.00019924 |  | [V]
Gain-calibration (SW1 & SW4 Closed)
Measure the (average) dc excited bridge output offset voltage with the ADC: | [12 empty] | GAINDC: |  | 0.99707 |  | [V/V]
ADC Codes (+ Polarity): |  | 083A415D | 083A47C5 | 083A494B | 083A4632 | 083A463E | 083A467D | 083A4795 | 083A44D9 | 083A4715 | 083A4694 | 083A43E0 | 083A4B2B | 083A45FB | 083A4983 | 083A41F7 | 083A4596
Voltages (+ Polarity): |  | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141
Measure the (average) ac excited bridge output offset voltages with the ADC: | [12 empty] | GAINAC: |  | 0.9971 |  | [V/V]
ADC Codes (+ Polarity): |  | 083A4772 | 083A48BE | 083A4BC5 | 083A4BFC | 083A47E3 | 083A46D9 | 083A465D | 083A4A73 | 083A4895 | 083A495E | 083A4BC0 | 083A4A4A | 083A4AED | 083A4BBC | 083A493F | 083A4C7F
Voltages (+ Polarity): |  | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141 | 0.010141
ADC Codes (- Polarity): |  | F7C575D6 | F7C577BE | F7C57A4B | F7C57A98 | F7C5780D | F7C575D7 | F7C57881 | F7C57A8B | F7C57843 | F7C57555 | F7C57694 | F7C574C4 | F7C579A9 | F7C575C6 | F7C57692 | F7C5761D
Voltages (- Polarity): |  | -0.010142 | -0.010142 | -0.010142 | -0.010142 | -0.010142 | -0.010142 | -0.010142 | -0.010142 | -0.010142 | -0.010142 | -0.010142 | -0.010142 | -0.010142 | -0.010142 | -0.010142 | -0.010142
Measure the (average) ac excited bridge output offset voltage with DMM1: | [12 empty] | GAINDMM: |  | 0.99229 |  | [V/V]
VFS_DMM_P: |  | 0.010188 | [V] |  | VREF_FS_DMM_P: |  | 5.0736 | [V] | → (Eq. 11) → |  | VFS_DMM_P_NORM: |  | 0.010189 | [V] | ↑ (Eq. 13) ↑
VFS_DMM_N: |  | -0.010195 | [V] |  | VREF_FS_DMM_N: |  | -5.0748 | [V] | → (Eq. 12) → |  | VFS_DMM_P_NORM: |  | -0.010192 | [V]
Bridge measurment error vs. temperature with 10 mV Input (SW1 & SW4 Closed)
Measure the (average) dc excited bridge output offset voltage with the ADC, for all desired temperature settings:
ADC Codes (+ Polarity, -40ᵒC): |  | 08072E29 | 080735FF | 08073156 | 08073800 | 08073BCB | 080738FC | 08073648 | 080734D7 | 080739FD | 08073965 | 08073BED | 080736F4 | 0807344A | 0807395F | 080732CC | 08073394
Voltages (+ Polarity, -40ᵒC): |  | 0.0098947 | 0.0098948 | 0.0098947 | 0.0098949 | 0.0098949 | 0.0098949 | 0.0098948 | 0.0098948 | 0.0098949 | 0.0098949 | 0.0098949 | 0.0098948 | 0.0098948 | 0.0098949 | 0.0098948 | 0.0098948
Average (+ Polarity, -40ᵒC): |  | 0.0098948 |  | [V] | → (Subtract OFFSETDC) & (Multiply by GAINDC) → |  |  |  |  | 0.0097549 |  | [V] | VFS_DC_-40°C |  | 0.0097549 |  | [V]
ADC Codes (+ Polarity, -20ᵒC): |  | 0819E0FB | 0819EFB1 | 0819E863 | 0819EE93 | 0819F096 | 0819EA9C | 0819EB3A | 0819F22E | 0819EBCA | 0819E6CD | 0819EE88 | 0819ECAB | 0819E7C3 | 0819E92B | 0819EAA7 | 0819EF03
Voltages (+ Polarity, -20ᵒC): |  | 0.0099847 | 0.009985 | 0.0099848 | 0.0099849 | 0.009985 | 0.0099849 | 0.0099849 | 0.009985 | 0.0099849 | 0.0099848 | 0.0099849 | 0.0099849 | 0.0099848 | 0.0099848 | 0.0099849 | 0.009985
Average (+ Polarity, -20ᵒC): |  | 0.0099849 |  | [V] | → (Subtract OFFSETDC) & (Multiply by GAINDC) → |  |  |  |  | 0.0098447 |  | [V] | VFS_DC_-20°C |  | 0.0098447 |  | [V]
ADC Codes (+ Polarity, 0ᵒC): |  | 0829BBFC | 0829C41E | 0829C296 | 0829BFF0 | 0829C05A | 0829C5F5 | 0829CA2D | 0829BD8C | 0829BDEC | 0829C059 | 0829C11D | 0829BF12 | 0829C107 | 0829BE23 | 0829BFC6 | 0829BBD3
Voltages (+ Polarity, 0ᵒC): |  | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061
Average (+ Polarity, 0ᵒC): |  | 0.010061 |  | [V] | → (Subtract OFFSETDC) & (Multiply by GAINDC) → |  |  |  |  | 0.0099207 |  | [V] | VFS_DC_0°C |  | 0.0099207 |  | [V]
ADC Codes (+ Polarity, 40ᵒC): |  | 0843A3C2 | 08439FA8 | 0843A0A3 | 08439D99 | 0843A470 | 0843A596 | 0843A33B | 0843A2EA | 0843A738 | 0843A3A2 | 0843A901 | 0843A53D | 0843A2E1 | 0843A2FB | 0843A1D1 | 0843AD3B
Voltages (+ Polarity, 40ᵒC): |  | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186
Average (+ Polarity, 40ᵒC): |  | 0.010186 |  | [V] | → (Subtract OFFSETDC) & (Multiply by GAINDC) → |  |  |  |  | 0.010045 |  | [V] | VFS_DC_40°C |  | 0.010045 |  | [V]
ADC Codes (+ Polarity, 80ᵒC): |  | 085A9523 | 085A9511 | 085AA3E7 | 085AAE2E | 085AA602 | 085AA849 | 085AA0EE | 085AA28B | 085AA7DB | 085AA169 | 085A9CE6 | 085AA266 | 085AA155 | 085AA29B | 085AA9F2 | 085A9C70
Voltages (+ Polarity, 80ᵒC): |  | 0.010296 | 0.010296 | 0.010296 | 0.010297 | 0.010297 | 0.010297 | 0.010296 | 0.010296 | 0.010297 | 0.010296 | 0.010296 | 0.010296 | 0.010296 | 0.010296 | 0.010297 | 0.010296
Average (+ Polarity, 80ᵒC): |  | 0.010296 |  | [V] | → (Subtract OFFSETDC) & (Multiply by GAINDC) → |  |  |  |  | 0.010155 |  | [V] | VFS_DC_80°C |  | 0.010155 |  | [V]
ADC Codes (+ Polarity, 100ᵒC): |  | 08643B1A | 08643B47 | 086448C0 | 086449E0 | 086451CC | 0864509F | 08644764 | 086450E8 | 086453D1 | 08644F3E | 08644CDC | 08646069 | 0864522C | 08644E70 | 0864532B | 08644FB2
Voltages (+ Polarity, 100ᵒC): |  | 0.010343 | 0.010343 | 0.010343 | 0.010343 | 0.010343 | 0.010343 | 0.010343 | 0.010343 | 0.010343 | 0.010343 | 0.010343 | 0.010343 | 0.010343 | 0.010343 | 0.010343 | 0.010343
Average (+ Polarity, 100ᵒC): |  | 0.010343 |  | [V] | → (Subtract OFFSETDC) & (Multiply by GAINDC) → |  |  |  |  | 0.010202 |  | [V] | VFS_DC_100°C |  | 0.010202 |  | [V]
ADC Codes (+ Polarity, 125ᵒC): |  | 087108BA | 0870FFB5 | 08711ED9 | 08710C45 | 0871107F | 087105CF | 08711564 | 08710850 | 087106B9 | 087107CA | 0871140F | 08710C7F | 08712150 | 08711909 | 087111B2 | 08710448
Voltages (+ Polarity, 125ᵒC): |  | 0.010404 | 0.010404 | 0.010405 | 0.010404 | 0.010404 | 0.010404 | 0.010405 | 0.010404 | 0.010404 | 0.010404 | 0.010405 | 0.010404 | 0.010405 | 0.010405 | 0.010404 | 0.010404
Average (+ Polarity, 125ᵒC): |  | 0.010404 |  | [V] | → (Subtract OFFSETDC) & (Multiply by GAINDC) → |  |  |  |  | 0.010263 |  | [V] | VFS_DC_125°C |  | 0.010263 |  | [V]
Measure the (average) ac excited bridge output offset voltage with the ADC, for all desired temperature settings:
ADC Codes (+ Polarity, -40ᵒC): |  | 08072D77 | 08072A51 | 08072D6E | 08073020 | 080728F6 | 08072B1B | 08072672 | 08072A4C | 08072954 | 08072B93 | 080730CA | 08072792 | 080728F5 | 08072506 | 080724DA | 08072B41
Voltages (+ Polarity, -40ᵒC): |  | 0.0098947 | 0.0098946 | 0.0098947 | 0.0098947 | 0.0098946 | 0.0098946 | 0.0098945 | 0.0098946 | 0.0098946 | 0.0098946 | 0.0098947 | 0.0098945 | 0.0098946 | 0.0098945 | 0.0098945 | 0.0098946
| ADC Codes (- Polarity, -40ᵒC): | F7F7E528 | F7F7E7B2 | F7F7E4F0 | F7F7E277 | F7F7E2D2 | F7F7E61B | F7F7E4FF | F7F7E824 | F7F7E565 | F7F7E8BD | F7F7E7C3 | F7F7ED55 | F7F7ECC1 | F7F7E880 | F7F7EA15 | F7F7E8EE
Voltages (- Polarity, -40ᵒC): |  | -0.0098991 | -0.0098991 | -0.0098991 | -0.0098992 | -0.0098992 | -0.0098991 | -0.0098991 | -0.0098991 | -0.0098991 | -0.0098991 | -0.0098991 | -0.009899 | -0.009899 | -0.0098991 | -0.009899 | -0.009899
Average (+ Polarity, -40ᵒC): |  | 0.0098946 |  | Average (- Polarity, -40ᵒC): |  |  | -0.0098991 |  | → (Eq. 2), (Subtract OFFSETDC), & (x GAINDC) → |  |  |  | VFS_AC_-40°C |  | 0.0097562 |  | [V]
ADC Codes (+ Polarity, -20ᵒC): |  | 0819DE34 | 0819DDBD | 0819E4F7 | 0819DC85 | 0819DD5B | 0819E4DA | 0819E53B | 0819E198 | 0819E25E | 0819E208 | 0819DFF9 | 0819E1D7 | 0819DFA2 | 0819DF31 | 0819DA40 | 0819DA05
Voltages (+ Polarity, -20ᵒC): |  | 0.0099846 | 0.0099846 | 0.0099848 | 0.0099846 | 0.0099846 | 0.0099848 | 0.0099848 | 0.0099847 | 0.0099847 | 0.0099847 | 0.0099847 | 0.0099847 | 0.0099847 | 0.0099847 | 0.0099846 | 0.0099846
| ADC Codes (- Polarity, -20ᵒC): | F7E56BEB | F7E56DA9 | F7E56B6B | F7E5693D | F7E56CB8 | F7E56C2B | F7E568CB | F7E569C5 | F7E5699D | F7E56B00 | F7E56A8D | F7E568BE | F7E569E5 | F7E567D2 | F7E56B65 | F7E56CE2
Voltages (- Polarity, -20ᵒC): |  | -0.0099881 | -0.009988 | -0.0099881 | -0.0099881 | -0.009988 | -0.0099881 | -0.0099881 | -0.0099881 | -0.0099881 | -0.0099881 | -0.0099881 | -0.0099881 | -0.0099881 | -0.0099881 | -0.0099881 | -0.009988
Average (+ Polarity, -20ᵒC): |  | 0.0099847 |  | Average (- Polarity, -20ᵒC): |  |  | -0.0099881 |  | → (Eq. 2), (Subtract OFFSETDC), & (x GAINDC) → |  |  |  | VFS_AC_-20°C |  | 0.0098455 |  | [V]
ADC Codes (+ Polarity, 0ᵒC): |  | 0829BD72 | 0829B96E | 0829BCA2 | 0829B8B6 | 0829BC6D | 0829B896 | 0829BF17 | 0829B719 | 0829BBE7 | 0829BF33 | 0829B868 | 0829BA2F | 0829BDE4 | 0829BCEC | 0829BAD1 | 0829B332
Voltages (+ Polarity, 0ᵒC): |  | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061 | 0.010061
| ADC Codes (- Polarity, 0ᵒC): | F7D5928E | F7D59115 | F7D591A9 | F7D59370 | F7D59171 | F7D593EF | F7D590F8 | F7D59303 | F7D591FD | F7D595DF | F7D59027 | F7D58F40 | F7D590DC | F7D59026 | F7D594A7 | F7D59229
Voltages (- Polarity, 0ᵒC): |  | -0.010064 | -0.010064 | -0.010064 | -0.010064 | -0.010064 | -0.010064 | -0.010064 | -0.010064 | -0.010064 | -0.010064 | -0.010064 | -0.010064 | -0.010064 | -0.010064 | -0.010064 | -0.010064
Average (+ Polarity, 0ᵒC): |  | 0.010061 |  | Average (- Polarity, 0ᵒC): |  |  | -0.010064 |  | → (Eq. 2), (Subtract OFFSETDC), & (x GAINDC) → |  |  |  | VFS_AC_-0°C |  | 0.0099216 |  | [V]
ADC Codes (+ Polarity, 40ᵒC): |  | 0843B01A | 0843B50E | 0843BD6C | 0843B2F7 | 0843B7DE | 0843B89F | 0843B2EB | 0843B1B6 | 0843B16E | 0843B02D | 0843B26E | 0843B1B3 | 0843B218 | 0843B0FD | 0843B3AE | 0843B3DF
Voltages (+ Polarity, 40ᵒC): |  | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186 | 0.010186
| ADC Codes (- Polarity, 40ᵒC): | F7BC8DE7 | F7BC92C0 | F7BC8C6E | F7BC906C | F7BC8F69 | F7BC9494 | F7BC928E | F7BC905C | F7BC8ABF | F7BC8FB5 | F7BC8EDE | F7BC8A2D | F7BC8FF1 | F7BC92F5 | F7BC8FEA | F7BC92B7
Voltages (- Polarity, 40ᵒC): |  | -0.010185 | -0.010185 | -0.010185 | -0.010185 | -0.010185 | -0.010185 | -0.010185 | -0.010185 | -0.010185 | -0.010185 | -0.010185 | -0.010185 | -0.010185 | -0.010185 | -0.010185 | -0.010185
Average (+ Polarity, 40ᵒC): |  | 0.010186 |  | Average (- Polarity, 40ᵒC): |  |  | -0.010185 |  | → (Eq. 2), (Subtract OFFSETDC), & (x GAINDC) → |  |  |  | VFS_AC_40°C |  | 0.010044 |  | [V]
ADC Codes (+ Polarity, 80ᵒC): |  | 085AB0DF | 085AB8B8 | 085AB0F8 | 085AB39B | 085ABB24 | 085AB705 | 085AB834 | 085AB900 | 085ABD3E | 085AB2DD | 085AB1E3 | 085ABFCC | 085AC059 | 085ABD75 | 085AB876 | 085AAD26
Voltages (+ Polarity, 80ᵒC): |  | 0.010297 | 0.010297 | 0.010297 | 0.010297 | 0.010297 | 0.010297 | 0.010297 | 0.010297 | 0.010297 | 0.010297 | 0.010297 | 0.010297 | 0.010297 | 0.010297 | 0.010297 | 0.010297
| ADC Codes (- Polarity, 80ᵒC): | F7A67330 | F7A67706 | F7A67E0B | F7A67BD8 | F7A67884 | F7A67273 | F7A67256 | F7A674F6 | F7A66EC7 | F7A66FE9 | F7A67C7C | F7A68169 | F7A67D84 | F7A66FC1 | F7A67469 | F7A67BF7
Voltages (- Polarity, 80ᵒC): |  | -0.010291 | -0.010291 | -0.010291 | -0.010291 | -0.010291 | -0.010291 | -0.010291 | -0.010291 | -0.010291 | -0.010291 | -0.010291 | -0.010291 | -0.010291 | -0.010291 | -0.010291 | -0.010291
Average (+ Polarity, 80ᵒC): |  | 0.010297 |  | Average (- Polarity, 80ᵒC): |  |  | -0.010291 |  | → (Eq. 2), (Subtract OFFSETDC), & (x GAINDC) → |  |  |  | VFS_AC_80°C |  | 0.010152 |  | [V]
ADC Codes (+ Polarity, 100ᵒC): |  | 0864890B | 08647B30 | 08648436 | 086475B7 | 08647BCD | 086460F2 | 08647509 | 08648362 | 08646FA6 | 086471AD | 08647BBD | 08647650 | 086486AA | 086482F2 | 08647B26 | 08648CCF
Voltages (+ Polarity, 100ᵒC): |  | 0.010344 | 0.010344 | 0.010344 | 0.010344 | 0.010344 | 0.010343 | 0.010344 | 0.010344 | 0.010344 | 0.010344 | 0.010344 | 0.010344 | 0.010344 | 0.010344 | 0.010344 | 0.010344
| ADC Codes (- Polarity, 100ᵒC): | F79D7920 | F79D7AFF | F79D7847 | F79D6F8B | F79D6383 | F79D70E6 | F79D6504 | F79D6F9A | F79D7BA0 | F79D66F6 | F79D684C | F79D71BF | F79D637A | F79D5CA5 | F79D4FB5 | F79D64D2
Voltages (- Polarity, 100ᵒC): |  | -0.010334 | -0.010334 | -0.010334 | -0.010335 | -0.010335 | -0.010335 | -0.010335 | -0.010335 | -0.010334 | -0.010335 | -0.010335 | -0.010335 | -0.010335 | -0.010335 | -0.010335 | -0.010335
Average (+ Polarity, 100ᵒC): |  | 0.010344 |  | Average (- Polarity, 100ᵒC): |  |  | -0.010335 |  | → (Eq. 2), (Subtract OFFSETDC), & (x GAINDC) → |  |  |  | VFS_AC_100°C |  | 0.010197 |  | [V]
ADC Codes (+ Polarity, 125ᵒC): |  | 08711049 | 08711EC8 | 08710D06 | 08710788 | 08711783 | 0871117C | 0871158E | 087110E6 | 0871245E | 08710860 | 0871056C | 08710E13 | 08711E40 | 0871008C | 08710F1C | 08710B47
Voltages (+ Polarity, 125ᵒC): |  | 0.010404 | 0.010405 | 0.010404 | 0.010404 | 0.010405 | 0.010404 | 0.010405 | 0.010404 | 0.010405 | 0.010404 | 0.010404 | 0.010404 | 0.010405 | 0.010404 | 0.010404 | 0.010404
| ADC Codes (- Polarity, 125ᵒC): | F7915D79 | F7915B73 | F791502A | F7916257 | F7916A00 | F79150F9 | F79166CA | F7915C58 | F791625C | F7915083 | F7915BBA | F79152FF | F7916024 | F791434F | F7915655 | F7915307
Voltages (- Polarity, 125ᵒC): |  | -0.010393 | -0.010393 | -0.010393 | -0.010393 | -0.010393 | -0.010393 | -0.010393 | -0.010393 | -0.010393 | -0.010393 | -0.010393 | -0.010393 | -0.010393 | -0.010393 | -0.010393 | -0.010393
Average (+ Polarity, 125ᵒC): |  | 0.010404 |  | Average (- Polarity, 125ᵒC): |  |  | -0.010393 |  | → (Eq. 2), (Subtract OFFSETDC), & (x GAINDC) → |  |  |  | VFS_AC_125°C |  | 0.010257 |  | [V]
Measure the (average) ac excited bridge output offset voltages with DMM1, for all desired temperature settings:
VFS_DMM_-40ᵒC_P: |  | 0.0099577 | [V] | VREF_FS_DMM_-40ᵒC_P: |  | 5.0789 | [V] | →(Eq. 5)→ | VFS_DMM_-40ᵒC_P_NORM: |  | 0.0099475 | [V] |  |  | ↓ (Eq. 7) ↓
VFS_DMM_-40ᵒC_N: |  | -0.0099554 | [V] | VREF_FS_DMM_-40ᵒC_N: |  | -5.0798 | [V] | →(Eq. 6)→ | VFS_DMM_-40ᵒC_P_NORM: |  | -0.0099432 | [V] | VFS_DMM_-40ᵒC: |  | 0.0097569 |  | [V]
VFS_DMM_-20ᵒC_P: |  | 0.010042 | [V] | VREF_FS_DMM_-20ᵒC_P: |  | 5.0771 | [V] | →(Eq. 5)→ | VFS_DMM_-20ᵒC_P_NORM: |  | 0.010035 | [V] |  |  | ↓ (Eq. 7) ↓
VFS_DMM_-20ᵒC_N: |  | -0.010044 | [V] | VREF_FS_DMM_-20ᵒC_N: |  | -5.0782 | [V] | →(Eq. 6)→ | VFS_DMM_-20ᵒC_P_NORM: |  | -0.010035 | [V] | VFS_DMM_-20ᵒC: |  | 0.0098458 |  | [V]
VFS_DMM_0ᵒC_P: |  | 0.010115 | [V] | VREF_FS_DMM_0ᵒC_P: |  | 5.0756 | [V] | →(Eq. 5)→ | VFS_DMM_0ᵒC_P_NORM: |  | 0.010111 | [V] |  |  | ↓ (Eq. 7) ↓
VFS_DMM_0ᵒC_N: |  | -0.010118 | [V] | VREF_FS_DMM_0ᵒC_N: |  | -5.0767 | [V] | →(Eq. 6)→ | VFS_DMM_0ᵒC_P_NORM: |  | -0.010112 | [V] | VFS_DMM_0ᵒC: |  | 0.0099218 |  | [V]
VFS_DMM_40ᵒC_P: |  | 0.010229 | [V] | VREF_FS_DMM_40ᵒC_P: |  | 5.0724 | [V] | →(Eq. 5)→ | VFS_DMM_40ᵒC_P_NORM: |  | 0.010231 | [V] |  |  | ↓ (Eq. 7) ↓
VFS_DMM_40ᵒC_N: |  | -0.010238 | [V] | VREF_FS_DMM_40ᵒC_N: |  | -5.0737 | [V] | →(Eq. 6)→ | VFS_DMM_40ᵒC_P_NORM: |  | -0.010238 | [V] | VFS_DMM_40ᵒC: |  | 0.010044 |  | [V]
VFS_DMM_80ᵒC_P: |  | 0.010326 | [V] | VREF_FS_DMM_80ᵒC_P: |  | 5.0686 | [V] | →(Eq. 5)→ | VFS_DMM_80ᵒC_P_NORM: |  | 0.010336 | [V] |  |  | ↓ (Eq. 7) ↓
VFS_DMM_80ᵒC_N: |  | -0.010344 | [V] | VREF_FS_DMM_80ᵒC_N: |  | -5.0701 | [V] | →(Eq. 6)→ | VFS_DMM_80ᵒC_P_NORM: |  | -0.010351 | [V] | VFS_DMM_80ᵒC: |  | 0.010152 |  | [V]
VFS_DMM_100ᵒC_P: |  | 0.010366 | [V] | VREF_FS_DMM_100ᵒC_P: |  | 5.0668 | [V] | →(Eq. 5)→ | VFS_DMM_100ᵒC_P_NORM: |  | 0.01038 | [V] |  |  | ↓ (Eq. 7) ↓
VFS_DMM_100ᵒC_N: |  | -0.010387 | [V] | VREF_FS_DMM_100ᵒC_N: |  | -5.0684 | [V] | →(Eq. 6)→ | VFS_DMM_100ᵒC_P_NORM: |  | -0.010398 | [V] | VFS_DMM_100ᵒC: |  | 0.010197 |  | [V]
VFS_DMM_125ᵒC_P: |  | 0.010417 | [V] | VREF_FS_DMM_125ᵒC_P: |  | 5.0644 | [V] | →(Eq. 5)→ | VFS_DMM_125ᵒC_P_NORM: |  | 0.010436 | [V] |  |  | ↓ (Eq. 7) ↓
VOS_DMM_125ᵒC_N: |  | -0.010444 | [V] | VREF_FS_DMM_125ᵒC_N: |  | -5.0661 | [V] | →(Eq. 6)→ | VFS_DMM_125ᵒC_P_NORM: |  | -0.01046 | [V] | VFS_DMM_125ᵒC: |  | 0.010256 |  | [V]